## Slide 1: Scintillator detectors with direct readout and MaPMTs
13. June 2017 | Ralf Engels

## Slide 2: Outline
Scope of SKADI
  Instrument detector requirements using SKADI instrument as target
Multi-Anode Photomultiplier (MaPMT)
Neutron Scintillator
Detector design & components
Conclusions

## Slide 3: SKADI Detector Requirements
Count-rate rate capability of up to 20 MHz@10% dead time
Detection efficiency of 80% @ 5Å with a gamma suppression of 10-5
Position resolution of 6 mm2, 3 mm2, 1 µs time resolution for TOF
Operation in vacuum, adaptable shape
Use of 2 detectors at different distance from sample for coverage of large q-range:
active area of each detector 100 x 100 cm2
first detector with 20 x 20 cm2 hole in the middle and small dead area around the hole

## Slide 4: Multianode Photomultiplier (MaPMT)
  Hamamatsu H8500 / H12700 :
  dense array of PMT channels with small amount of dead space (~89 % active area)
  possibility of modular and scalable design for large areas
  High-Resolution Anger Mode
    position resolution of ~0.6 - 0.7 mm
  also: High-Rate Mode (SoNDe):
    count rate capability theoretically up to max. 1 MHz per MaPMT
Figures from Hamamatsu

## Slide 5: Hamamatsu H8500 / H9500 Flat-Panel MaPMT
Common Properties:
  Outer dimension 52 mm, active area 49 mm, peak wavelength 400 nm
  Gain ~ 106, but inhomogeneous of factor 2-3 between different pixels
  Maximum pulsed anode current >100 μA possible, but permanent current should be considerably lower for longer lifetime (~20 μA)
Figures from Hamamatsu
  H8500:
  Array of 8x8 pixel
  pitch 6.08 mm
  3% crosstalk
  H9500:
  Array of 16x16 pixel
  pitch 3.04 mm
  5% crosstalk

## Slide 6
Standard Neutron Scintillator:
GS20-6Li-glass
Neutron capture reaction:
6.6 weight% Li, 95% 6Li-enriched
Emission peak at ~390 nm (Ce doped)
Light yield ~ 6000 photons/n (corresponds ~1.5 MeV gamma)
Neutronen Sensitivity
    65% @ 1 Å
    85% @ 2 Å
www.scintacor.com
Scintillator with support glass
Pulse height spectra

## Slide 7: Readout Electronics Evaluation System
Required parameter for pulse processing: charge per neutron
  dependent on number of photons hitting photocathode, quantum efficiency, gain of MaPMT
  difficult to predict (reflective effects)
  adjustable to some degree by MaPMT gain
Pictures of the ROSMAP
Readout System attached to a MaPMT

## Slide 8: Readout Electronics Evaluation System
ROSMAP readout system for evaluation
  digitization and counting mode available, but only digitization mode used for tests
  2x VA32HDR14.3 ASICs for digitization of channels with 10:1 (changed to 3:1) charge splitter for measurement up to 200pC input charge
  trigger derived from PMT dynode signal
  14 bit ADC, data values are delivered with 8 bit resolution via ModBus interface
  read out rate of ~50 Hz achieved for digitization mode (with python interface)
  external high voltage supply for MaPMT

## Slide 9: System Overview

## Slide 10: 3D Model of one MaPMT Module and the Trenz-Lifter
ASIC Board with Aluminum housing
Controller / Backend Board
Trenz TE0720
Trenz – Lifter to remove safely the board from the ASIC board

## Slide 11: Test Setup @ TREFF
Picture of H8500 with the adapter
board and blue LED
View into the direction of the sample

## Slide 12: Conclusion
Neutron Scintillation Detector based on H8500 MaPMT
  Large active area → scalable design for larger detector size
  Higher resolution due to smaller pitch size of neighboring light detectors (~ 6mm)
  Higher count rate capability with smaller readout units (higher segmentation, smaller neutron occupancy)
  First prototype for evaluation based on ROSMAP system
  Images acquired with ROSMAP system look reasonable
Prototype measurements
  Pixel n/γ –discrimination is possible, but need more test for optimisation
  Backend Board was designed
  Adaption to the ASIC board was successfully
  Currently under commissioning

## Slide 13
Thanks !