## Slide 1: WP 9.2.1 Development of a WSF detector at ISIS for reflectometry applications
Jeff Sykora, Erik Schooneveld, Nigel Rhodes, Martin Sharkey, Sarah Mann and Nick Ferguson
28 May 2019
Bilbao

## Slide 2: Work Package 9.2.1 Objectives and Deliverables
Spatial resolution of 1-3mm, with a perspective of improvement into the sub-mm regime.
Time resolution better than 100μs.
Local instantaneous rate capability of several kHz/mm2.
D1. Initial WLS fibre detector hardware
D2. Interim report on scintillation detector development programme
D3. Final report on scintillation detector development programme

## Slide 3: Overview
Recap
  SHARD2 successes
  SHARD2 problems
New WSF detector concepts
A solution for INTER

## Slide 4: SHARD2 Configuration
1 mm fibres – 16 mm wide
Individual segments
  Optically divided
  Supports
    3D printed
  Strips of scintillator front and back
Single 64 ch FP-PMT per segment
Neutron beam

## Slide 5: SHARD2 Testing
OffSpec reflectometer
  Efficiency
  Rate capability
  Ghosting

## Slide 6: OffSpec Results
Detector in direct beam
  Measure efficiency
  Global and local peak rates
80mm
65mm
16mm

## Slide 7: SHARD2 Results
Detector in direct beam
  Measure efficiency
  Global and local peak rates
80mm
65mm
16mm
225 μm front scintillator + 450 μm back scintillator is ~90% efficient at 9 Å
450 μm front scintillator efficiency drops to 45% at 9 Å
“Local” (within beam area) instantaneous peak rates
  Linear until ~160 kcps

## Slide 8: SHARD2 Reflectivity Results
3He point detector
SHARD inner 2 segments
Polymer on Si substrate
Exact match with similar error bars despite no shielding and less than ½ the active area!

## Slide 9: Ghosting
Si reflectivity
Transmitted
Refracted
Reflected

## Slide 10: Ghosting
Ghosting is an issue with this coding scheme
  N.B. This was expected
Example:
Fibre no. | PMT 1 | PMT 2
1 | 1 | 9
2 | 1 | 10
3 | 2 | 10
4 | 2 | 11
5 | 3 | 11
6 | 3 | 12
7 | 4 | 12
8 | 4 | 13
9 | 5 | 13
10 | 5 | 14
11 | 6 | 14
12 | 6 | 15
13 | 7 | 15
14 | 7 | 16
15 | 8 | 16
16 | 8 | 9
17 | 3 | 9
18 | 3 | 10
19 | 4 | 10
20 | 4 | 11
21 | 5 | 11
22 | 5 | 12
23 | 6 | 12
24 | 6 | 13
25 | 7 | 13
26 | 7 | 14
27 | 8 | 14
28 | 8 | 15
29 | 1 | 15
30 | 1 | 16
31 | 2 | 16
32 | 2 | 9
Direct beam
Location of ghosts
  Ghosts caused by :
  Afterglow
  Coding
  Cross-talk

## Slide 11: New geometry
Takes advantage of reflectometer operation
Optimise rate capability in one section
Eliminates ghosting
Remains cost-effective

## Slide 12: Optimisation of rate capability and Ghost elimination
Uncoded fibres in detectors pixels seeing highest rate – specular reflection
  Off specular is very low rate
Optical dividers between fibres
Incident
beam ki
Transmitted
beam
Reflected
beam kf
qz
qy
qx
θ1

## Slide 13: Reflectometer Operation
Full q reconstruction
  Time of flight
  Several angles to extend q range
Reflected beam on same detector pixels
Incident
beam ki
Transmitted
beam
Reflected
beam kf
qz
qy
qx
Incident
beam ki
Transmitted
beam
Reflected
beam kf
qz
qy
qx
θ1
Δz
θ2

## Slide 14: Layout of detector with dividers
Flat
Flat
w/dividers
0.25mm scintillator
0.45mm scintillator
Foil reflector
B4C backstop
Steel dividers
Neutrons

## Slide 15: Fibre holder and dividers
Machined Al block
  Individual holes for fibres
  Pocket to contain B4C and scintillator
Stainless steel dividers
  Sub-100 μm thickness
  Photo-fabricated
  Contains light spread

## Slide 16: Effect of dividers Efficiency
~5% gain in efficiency because of dividers
  Removes ambiguity in event determination
Flat
Flat
w/dividers

## Slide 17: Light spread Average intensity distribution across PMT
Flat without dividers
38%
11%
10%
34%
3%
PMT 21
PMT 22
PMT 6
PMT 7
PMT 23

## Slide 18: Light spread Average intensity distribution across PMT
Dividers reduce light spread by a factor 2
Flat without dividers
Flat with dividers
38%
17%
18%
3%
2%
11%
10%
34%
3%
4%
PMT 61
PMT 62
PMT 46
PMT 47
PMT 63

## Slide 19: Light spread Average intensity distribution across PMT
Dividers reduce light spread by a factor 2
  It is worth putting the dividers in unless it hinders ghost rejection
Flat without dividers
Flat with dividers
38%
17%
20%
5%
2%
12%
10%
34%
3%
5%
PMT 32
PMT 21
PMT 5
PMT 6
PMT 22

## Slide 20: Light spread Average intensity distribution across PMT
Dividers reduce light spread by a factor 2
  It is worth putting the dividers in unless it hinders ghost rejection
Flat without dividers
Flat with dividers
38%
17%
20%
5%
2%
12%
10%
34%
3%
5%
PMT 64
PMT 53
PMT 37
PMT 38
PMT 54

## Slide 21: Optimisation of rate capability and Ghost elimination
Uncoded fibres in detectors pixels seeing highest rate – specular reflection
  Off specular is very low rate
Optical dividers between fibres
Tested on INTER
Incident
beam ki
Transmitted
beam
Reflected
beam kf
qz
qy
qx
θ1

## Slide 22: Ghost elimination Success

## Slide 23: 0.6 mm demonstrator
0.5mm fibres
0.1mm dividers
High rate walking coincidence (no repeated codes)

## Slide 24: First characterisation of 0.6 mm
Verifying position resolution on EMMA

## Slide 25: Verifying position resolution
EMMA
  Resolution
    0.2mm slit

## Slide 26: Verifying position resolution
Integrated in time of flight
Single pixel illumination
Resolution = 0.6 mm

## Slide 27: Verifying position resolution
Low scattering from fibres and B4C behind scintillator
Beam spread dominated by divergence
No exact numbers yet
  Good enough for INTER

## Slide 28: Detector spec summary
Requirements laid out in WP 9.2 have been satisfied
First application is the INTER beamline
 | Requirements | SHARD2
Neutron Detection Efficiency | 40% at 1.8 A | 45% (0.25mm front scint) 60% (0.45mm front scint) at 1.8 Å
Gamma Sensitivity | 10-6 at 1MeV | 3 x 10-7 at 1MeV
Position Resolution | 1-3 mm | 0.6 mm
Timing Resolution | 100 μs | 10 μs
Rate Capability | (several) kcps/mm2 | 1 kcps/mm2
Placement Accuracy | 99.99% | 99.99%

## Slide 29: First Application: INTER
Installing new guide
  Focussing guide
  Increase in flux spread out along the detector
Current
New guide

## Slide 30: First Application: INTER
Installing new guide
  Focussing guide
  Increase in flux spread out along the detector
Replace old linear detector
  New SHARD style to cover the area and match count rate requirements
  Temp replacement = 0.6mm segment

## Slide 31: 0.6mm segment on INTER
Rates
  16 kcps per 30mm2 area
D2O reflectivity at low angle
Divergent mode on 3He tube
40mm x 10mm slits

## Slide 32: Early results on INTER
Detector covers full divergent mode in vertical plane
Resolution and uniformity showed existing guide defects
In operation this user cycle
Highly fringed sample in divergent mode

## Slide 33: Mid-term detector design for INTER
8 Segments – 64 channel PMT per segment
150 μm gap between segments
2 different widths to match beam profile

## Slide 34: Summary
A sub-mm position resolution detector has been developed for reflectometry
A new high rate section has been successfully implemented keeping the detector cost-efective
So far a detector capable of 1 kcps/mm2has been built
Ghosting has been eliminated
A new solution for the INTER reflectometer is now being implemented due to these developments.

## Slide 35: Thank You!

## Slide 36: Other issues?

## Slide 37: Conclusions
Neutron counts are “missing” in the un-coded region
  Significant portion of events are unidentified
  Unidentified are in neighbouring fibres
    Most likely neutrons absorbed on the boundary between two fibres
PMT 3
PMT 4
PMT 1
PMT 2
PMT 3
PMT 4
PMT 1
PMT 2
PMT 5

## Slide 38: Conclusions
Neutron counts are “missing” in the un-coded region
  Significant portion of events are unidentified
  Unidentified are in neighbouring fibres
    Most likely neutrons absorbed on the boundary between two fibres
No multi-counting in the coded areas
PMT 3
PMT 4
PMT 1
PMT 2
PMT 3
PMT 4
PMT 1
PMT 2
PMT 5

## Slide 39: EMMA testing Extra “bumps”